Unit: IBT												Date: August 2017

Participants of the risk assessment: Torun M. Melø (senior engineer)

Short description of activity: The use of NMR instruments and filling of cryogenic liquids into the magnets

| ID nr | Activity/process | Responsible person | Laws, regulations etc. | Existing documentation | Existing safety measures | Comment | Can the task be carried out alone? |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | In general when working with cyrogenic liquids |  |  |  | Never use wellies or boots. The proper shoe wear should be tight fitting and easily removable. |  |  |
| 1 | Filling of liquid nitrogen into tank | Torun M. Melø |  |  | Use protective clothing; covering clothes, leather gloves and goggels.An O2 sensor har been installed in the room | A portabel O2 sensor is recommended and JALA | Yes |
| 2 | Liquid nitrogen filling of magnet vessel | Torun M. Melø |  | Safety note from supplier | Use protective clothing; covering clothes, leather gloves and goggels. An O2 sensor has been installed in the room. | A portabel O2 sensor is recommended and JALA | Yes |
| 3 | Liquid helium filling og magnet vessel | Torun M. Melø |  | Safety notes from supplier | Use protective clothing; lab coat, leather gloves and goggels. Always done by two persons. |  | No |
| 4 | Common use of the NMR instruments | Torun M. Melø |  | Safety note from supplier | Warning sign placed on the door. The operation of medical electronic implants, such as pacemakers, may be affected by the magnetic field. Other medical implants, may contain ferromagnetic material and would therefore be subjected to strong attractive forces near the magnet.All user will be given training. |  | Yes |
| 5 | Investigation and repair of electrical units | Torun M. Melø |  | Safety note from supplier | When testing and repairing electrical units, the unit itself has to be turned off and unplugged | There should be two persons present when carrying out this task | No |

Unit: 	IBT									Date: August 2017

Line manger: Kjetil Rasmussen

Participant Torun M. Melø, (senior engineer) resposable for the NMR-lab

| ID nr | Activity from the identification process form | Potential undesirable incident/strain | Likelihood: | Consequence: | Risk value | Comments/status Suggested measures |
| --- | --- | --- | --- | --- | --- | --- |
| ID Nr |  |  | Likelihood (1-5) | Human (A-E) | Environment (A-E) | Economy/ material (A-E) | Reputation (A-E) |  |  |
| 1 | Filling of liquid nitrogen into tank | Spillage of liquid N2 can lead to reduced amount of O2 in air. | 2 | C | A | A | C | 2A-C | The filling under a fume point (spot exhauster). An O2 sensor has been installed in the room. |
|  |  | Skin contact with liquid N2 can lead to cold burns. | 3 | B | A | A | B | 3A-B | Use skin covering clothing;, leather gloves and goggels. |
| 2 | Liquid nitrogen filling of magnet vessel | Magnet quench. | 1 | A | A | D | A | 1A-D |  |
|  |  | Condenstion of O2 can lead to explotion or fire. | 1 | C | A | D | C | 1A-D | Avoid open flames and oil/fat spills in the immediate vicinity of the magnet system. |
|  |  | Skin contact with liquid N2 can lead to cold burns. | 2 | A | A | A | A | 2A | Use protective clothing; lab coat, leather gloves and goggels. |
|  |  | Spillage of liquid N2 can lead to reduced amount of O2 in air | 1 | C | A | A | C | 1A-C | The use of a portabel O2 sensor is recommended. |
| 3 | Liquid helium filling og magnet vessel | Magnet quench | 2 | A | A | D | A | 2A-D |  |
|  |  | Skin contact with liquid He can lead to cold burns. | 2 | A | A | A | A | 2A | Use skin covering clothing; leather gloves and goggels. |
|  |  | Condenstion of O2 can lead to explotion or fire. | 1 | C | A | D | C | 1A-D | Avoid open flames and oil/fat spills in the immediate vicinity of the magnet system. |
| 4 | Common use of the NMR instruments | Ferromagnetic devices get stuck to the magnet | 1 | A | A | D | A | 1A-D | Empty pockets befor approching magnet |
|  |  | Accidental breaking of NMR test tubes can lead to spillage of potentially harmful chemicals. | 3/4 | B | A | B | A | 3/4A-B | The lab should be equipped with absorbant pad (pig hazmat) for cleaning of chemical spillage. |
|  |  | The solvent chloroform will evaporate from NMR test tubes, even when using caps on the tubes. Therefore, all NMR test tubes, some of which contain chloroform, are to be stored in the fume hood, after being run in the autosampler | 3 | A | A | A | A | 3A | These NMR tubes are place under a fume hood. |
|  |  | Errors done by the user. (e.g. forgetting to use a spinner, putting the depth gauge in the magnet, the use of harmful acquisition parameters) | 3 | A | A | A | A | 3A |  |

# Risk assessment matrix NTNU

| CONSEQUENCE | Very serious | E1 | E2 | E3 | E4 | E5 |
| --- | --- | --- | --- | --- | --- | --- |
|  | Serious | D1 | D2 | D3 | D4 | D5 |
|  | Moderate | C1 | C2 | C3 | C4 | C5 |
|  | Small | B1 | B2 | B3 | B4 | B5 |
|  | Very small | A1 | A2 | A3 | A4 | A5 |
|  |  | Very small | Small | Moderate | High | Very high |
|  |  | PROBABILITY |

Color-coding in the risk assessment matrix

| Color | Description |
| --- | --- |
| Red |  | Unacceptable risk. Measures must be implemented to reduce risk. |
| Yellow |  | Evaluation area. Evaluate whether measures need to be implemented. |
| Green |  | Acceptable risk. Measures may be considered based on other factors. |